## Slide 1
INSIDE EMBRYO
BY AURION
Micromanipulator
Setup
Building a reliable IVF workstation
A beginner training deck on readiness, control and safe escalation
BEGINNER LEVEL
Prepared by Manoj Kumar K, Embryologist
Inside Embryo by Aurion • 9 August 2026
Educational illustration; not a clinical photograph.
Notes: Open by framing setup as a release decision, not as a collection of knobs. The deck is educational and must be paired with supervised training and the local SOP.
[Sources]
- https://doi.org/10.1093/humrep/deag096
- https://doi.org/10.1093/hropen/hoad001

## Slide 2
Four decisions define a safe beginner setup
The goal is reasoning, not memorising device-specific values
02
4
By the end, you should be able to…
1
SEE
Identify the optical path, manipulators, injectors, heated stage and pipettes.
2
CONTROL
Explain how movement, focus, pressure and environment interact.
3
RELEASE
Follow a supervised pre-use sequence and functional check.
4
STOP
Recognise faults that require correction, rechecking or escalation.
This deck does not authorise unsupervised manipulation of human gametes or embryos.
INSIDE EMBRYO BY AURION • EMBRYOLOGY & IVF LABORATORY SCIENCE
MANOJ KUMAR K • 02
Notes: Use the four verbs to establish the learning contract. Emphasise that competence requires supervised practice, observed performance and local authorisation.
[Sources]
- https://doi.org/10.1093/humrep/deag096
- https://doi.org/10.1093/hropen/hoad001

## Slide 3
Setup is a readiness gate in the ART pathway
It prepares a platform; it does not decide the treatment indication
03
1
Clinical plan
Indication, consent, identity
2
Gamete prep
Oocyte and sperm readiness
3
Station release
Setup + functional checks
4
Procedure
ICSI, hatching or biopsy
5
Post-procedure
Assessment, culture, traceability
Clinical choice and technical readiness answer different questions
A workstation can be technically ready while a procedure is not clinically indicated—and a valid indication cannot compensate for an unready station.
INSIDE EMBRYO BY AURION • EMBRYOLOGY & IVF LABORATORY SCIENCE
MANOJ KUMAR K • 03
Notes: Separate two decisions: whether a procedure should be used and whether the station is technically fit for use. This protects against treating equipment capability as clinical justification.
[Sources]
- https://doi.org/10.1093/humrep/deag029
- https://www.asrm.org/practice-guidance/practice-committee-documents/intracytoplasmic-sperm-injection-for-nonmale-factor-indications-a-committee-opinion-2026/
- https://doi.org/10.1093/humrep/deag096

## Slide 4
Four interfaces define reliable setup
A fault in one interface can invalidate the whole release decision
04
SETUP
CONTROL
VISION
focus • contrast • focal plane
MOTION
coarse travel • fine control • drift
PRESSURE
aspiration • neutral • expulsion
ENVIRONMENT
temperature • exposure • vibration
Optical clarity without stable movement is not readiness.
Precise movement without predictable pressure is not readiness.
INSIDE EMBRYO BY AURION • EMBRYOLOGY & IVF LABORATORY SCIENCE
MANOJ KUMAR K • 04
Notes: Use this as the deck's organising model. The four interfaces are a teaching abstraction; the local equipment qualification and SOP determine the exact checks.
[Sources]
- https://doi.org/10.1093/humrep/deag096

## Slide 5
The workstation behaves as one system
Each component has a distinct control job
05
1
Inverted microscope
Makes the dish and pipette tips visible.
2
Manipulators
Translate hand movement into controlled three-axis travel.
3
Microinjectors
Create controlled negative, neutral or positive pressure.
4
Heated stage
Supports validated handling temperature at the working surface.
5
Pipette holders
Secure geometry and connect the pressure path.
6
Anti-vibration support
Reduces external mechanical disturbance.
1
2
3
4
5
6
Educational illustration; device layouts vary.
INSIDE EMBRYO BY AURION • EMBRYOLOGY & IVF LABORATORY SCIENCE
MANOJ KUMAR K • 05
Notes: Point out that local device design varies. The conceptual control jobs remain: see, move, apply pressure, protect the environment and stabilise the platform.
[Sources]
- https://doi.org/10.1093/humrep/deag096

## Slide 6
Two arms, two jobs, one focal plane
Holding creates stability; the working side creates the intended procedure-specific action
06
H
W
HOLDING SIDE
WORKING SIDE
✓
Blunt tip stabilises the cell or embryo.
✓
Negative pressure is gentle and controlled.
✓
Excess suction or poor centring can distort the target.
✓
Fine, procedure-specific pipette performs the action.
✓
Geometry, bevel and diameter must match the validated task.
✓
Blocked, damaged or misaligned tips require replacement.
ONE
FOCAL
PLANE
INSIDE EMBRYO BY AURION • EMBRYOLOGY & IVF LABORATORY SCIENCE
MANOJ KUMAR K • 06
Notes: Do not teach a universal left/right convention; systems and local workflows vary. Teach the functional distinction between stabilising and working sides.
[Sources]
- https://doi.org/10.1093/humrep/deag029
- https://doi.org/10.1093/humrep/deag096

## Slide 7
Alignment is three-dimensional
Both tips must be centred, sharply focused and mechanically stable
07
Educational illustration; no penetration is shown.
THE 3-POINT ALIGNMENT TEST
1
Same focal plane
Both tips remain sharp without refocusing.
2
Coaxial approach
Tips share a stable horizontal working axis.
3
Clear travel
Movement is smooth, centred and free of dish contact.
Any change = STOP and realign.
INSIDE EMBRYO BY AURION • EMBRYOLOGY & IVF LABORATORY SCIENCE
MANOJ KUMAR K • 07
Notes: Alignment is checked before introducing reproductive material. Avoid teaching by appearance alone; stability through fine movement matters.
[Sources]
- https://doi.org/10.1093/humrep/deag096
- https://doi.org/10.1093/hropen/hoaa020

## Slide 8
Pressure must be predictable before use
Test response in acellular droplets—not on reproductive material
08
NEGATIVE
←
ASPIRATION / HOLD
Controlled inward flow
NEUTRAL
•
NO INTENDED FLOW
No spontaneous drift
POSITIVE
→
EXPULSION
Controlled outward flow
Leaks, bubbles or blockage make pressure delayed or non-linear.
Use the manufacturer instructions, validated local SOP and supervised training—never a universal memorised setting.
INSIDE EMBRYO BY AURION • EMBRYOLOGY & IVF LABORATORY SCIENCE
MANOJ KUMAR K • 08
Notes: Describe direction and response, not numeric settings. The acceptable response is device-, pipette- and SOP-specific and should be functionally checked before use.
[Sources]
- https://doi.org/10.1093/humrep/deag096
- https://doi.org/10.1093/hropen/hoaa020

## Slide 9
Microenvironment instability blocks release
Setup must protect reproductive material during the handling interval
09
TEMPERATURE
Risk: heat loss at the dish and working surface
Control: verify the heated stage with a calibrated method
pH & EXPOSURE
Risk: prolonged handling or unsuitable buffering
Control: validated media, timing and local gas/air workflow
LIGHT
Risk: unnecessary optical exposure
Control: adequate visibility with minimal exposure
VIBRATION & CLEANLINESS
Risk: tip drift, mechanical shock or contamination
Control: stable support, clean surfaces and controlled access
Do not copy a numerical target from a teaching slide: use the validated range in the local SOP.
INSIDE EMBRYO BY AURION • EMBRYOLOGY & IVF LABORATORY SCIENCE
MANOJ KUMAR K • 09
Notes: ESHRE recommends appropriate heated devices, calibrated measurement, accepted ranges and correction of out-of-range parameters. The local validation defines the release range.
[Sources]
- https://doi.org/10.1093/humrep/deag096

## Slide 10
Eight steps create a release decision
The sequence is conceptual; the local SOP supplies exact checks and ranges
10
1
Review
SOP, procedure, readiness criteria
2
Stabilise
Clean, power and allow station stability
3
Verify
Heated-stage temperature
4
Mount
Correct pipettes and secure holders
5
Prime
Pressure path; no bubbles or leaks
6
Align
Tips in one focal plane
7
Test
Movement and pressure acellularly
8
Release
Document or escalate
INSIDE EMBRYO BY AURION • EMBRYOLOGY & IVF LABORATORY SCIENCE
MANOJ KUMAR K • 10
Notes: Walk through the readiness path. If any critical check fails, the pathway stops until the fault is corrected, rechecked and documented or escalated.
[Sources]
- https://doi.org/10.1093/humrep/deag096
- https://doi.org/10.1093/hropen/hoaa020

## Slide 11
Functional checks authorise release
Perform the checks without reproductive material in the dish
11
TEST IN ACELLULAR DROPLETS
✓
Movement
Coarse and fine travel are smooth and predictable.
✓
Focus
Both tips remain sharp in the intended focal plane.
✓
Pressure
Negative, neutral and positive response are controlled.
✓
Stability
No obvious leak, bubble, vibration or spontaneous drift.
✓
Environment
Validated temperature and handling conditions are in range.
RELEASE?
✓
YES
All critical checks pass
and the record is complete
Otherwise: HOLD • CORRECT • RECHECK
INSIDE EMBRYO BY AURION • EMBRYOLOGY & IVF LABORATORY SCIENCE
MANOJ KUMAR K • 11
Notes: A station should be released only when every critical criterion passes and the check is recorded. An unrecorded check is not a complete release process.
[Sources]
- https://doi.org/10.1093/humrep/deag096

## Slide 12
Stop-use signals override the schedule
Protect the material first; then correct, recheck and document
12
STOP
!
TIP DRIFT
Movement continues after input stops
!
BUBBLE / LEAK
Delayed or non-linear pressure response
!
DAMAGED TIP
Cracked, blocked or unstable mount
!
TEMP. OUT
Working surface outside validated range
!
VIBRATION
Visible disturbance or stage movement
!
IDENTITY GAP
Dish, witness or documentation uncertainty
INSIDE EMBRYO BY AURION • EMBRYOLOGY & IVF LABORATORY SCIENCE
MANOJ KUMAR K • 12
Notes: Stop-use criteria should be defined by the laboratory and equipment SOPs. Faulty equipment should be clearly labelled out of use to prevent accidental use.
[Sources]
- https://doi.org/10.1093/humrep/deag096

## Slide 13
Troubleshoot the symptom first
Use the local fault tree and escalate when the cause is uncertain
13
OBSERVED SYMPTOM
POSSIBLE CAUSE
SAFE FIRST RESPONSE
RELEASE CONDITION
Tip drifts
Loose mount, mechanical play, vibration or thermal instability
Remove live dish; secure, stabilise and retest
No drift through fine movement
Pressure is delayed
Bubble, leak, blocked tip or loose connection
Inspect path; reprime or replace per SOP
Predictable acellular response
Tips will not share focus
Height or holder angle mismatch
Re-establish geometry from coarse to fine
Both tips sharp in one plane
Stage reading is out
Warm-up incomplete, probe issue or device fault
Hold use; verify with calibrated method
Validated range restored and recorded
Never troubleshoot a persistent fault with reproductive material still on the station.
INSIDE EMBRYO BY AURION • EMBRYOLOGY & IVF LABORATORY SCIENCE
MANOJ KUMAR K • 13
Notes: These are conceptual examples, not a substitute for a device-specific fault tree. Persistent or unexplained faults require escalation and an out-of-use label.
[Sources]
- https://doi.org/10.1093/humrep/deag096

## Slide 14
Shared platform, procedure-specific setup
Change pipettes, accessories and competence requirements only through the validated workflow
14
ICSI
CORE TOOLS
Holding + injection pipette
EXTRA CONTROL
Fine pressure control and oocyte orientation
BEGINNER LIMIT
Indication is a clinical decision
ASSISTED HATCHING
CORE TOOLS
Holding pipette + validated zona-opening method
EXTRA CONTROL
Optical alignment; laser or mechanical system if authorised
BEGINNER LIMIT
Not universally required
EMBRYO BIOPSY
CORE TOOLS
Holding + biopsy pipette; laser where applicable
EXTRA CONTROL
Cell removal, tubing and traceability workflow
BEGINNER LIMIT
Requires specific training and SOP
Never transfer settings between procedures or devices without validation, competency and SOP control.
INSIDE EMBRYO BY AURION • EMBRYOLOGY & IVF LABORATORY SCIENCE
MANOJ KUMAR K • 14
Notes: Emphasise the shared control principles while avoiding the idea that one setup automatically qualifies a different procedure. Biopsy recommendations require trained practitioners and SOP adherence.
[Sources]
- https://doi.org/10.1093/humrep/deag029
- https://doi.org/10.1093/hropen/hoaa020
- https://www.asrm.org/practice-guidance/practice-committee-documents/intracytoplasmic-sperm-injection-for-nonmale-factor-indications-a-committee-opinion-2026/

## Slide 15
Scenario: the working pipette drifts
What should a supervised trainee do next?
15
OBSERVATION
The tip continues to move after input stops; reproductive material is nearby.
1
PAUSE & PROTECT
Stop manipulation; secure or remove the live dish according to the SOP.
2
TEST ACELLULARLY
Confirm drift without reproductive material on the station.
3
CHECK THE SYSTEM
Mount, mechanical play, vibration, stage stability and pressure path.
4
RETEST & DECIDE
Stable response and completed record—or out-of-use label and escalation.
Correct answer: protect the material first. Troubleshooting comes second.
INSIDE EMBRYO BY AURION • EMBRYOLOGY & IVF LABORATORY SCIENCE
MANOJ KUMAR K • 15
Notes: Invite the audience to answer before revealing the pathway. The correct first action is to protect the reproductive material and stop use; do not continue while trying to compensate manually.
[Sources]
- https://doi.org/10.1093/humrep/deag096

## Slide 16
Reliable setup belongs to the quality system
Daily checks cannot replace validation, maintenance or competency
16
VALIDATION / QUALIFICATION
Design, installation, operational and performance evidence
CALIBRATION & MAINTENANCE
Calibrated measurement and documented preventive work
SOP & CHANGE CONTROL
Current instructions; controlled changes and deviations
COMPETENCY
Supervised training, observation and ongoing performance review
DAILY RELEASE RECORD
Pre-use checks, fault action, recheck and authorisation
A COMPLETE RELEASE RECORD
✓
Date, station and procedure
✓
Operator and witness where required
✓
Temperature and functional results
✓
Pipette / consumable traceability
✓
Fault, correction and recheck
✓
Release or escalation decision
No record = no completed release.
INSIDE EMBRYO BY AURION • EMBRYOLOGY & IVF LABORATORY SCIENCE
MANOJ KUMAR K • 16
Notes: Relate daily readiness to the broader quality system. ESHRE describes equipment qualification, calibrated measurement, maintenance, SOP control, operator assessment and performance monitoring.
[Sources]
- https://doi.org/10.1093/humrep/deag096
- https://doi.org/10.1093/hropen/hox011
- https://doi.org/10.1093/hropen/hoad001

## Slide 17
Technical readiness is not a clinical outcome
Use the correct language: observation, association, prediction, diagnosis and guarantee are not interchangeable
17
READINESS CAN SHOW
READINESS CANNOT SHOW
✓
Tips are aligned and stable
✓
Movement is controlled
✓
Pressure response is predictable
✓
Validated environment is in range
✓
Checks and release are documented
×
Whether an oocyte is biologically competent
×
Chromosome status or embryo potential
×
Fertilisation, implantation or pregnancy
×
Live birth for an individual patient
×
Whether a procedure is clinically indicated
≠
A ready station reduces avoidable technical risk; it cannot guarantee an ART outcome.
INSIDE EMBRYO BY AURION • EMBRYOLOGY & IVF LABORATORY SCIENCE
MANOJ KUMAR K • 17
Notes: This is the central safety distinction. Technical observations support process control, but they are not diagnostic or predictive of an individual treatment outcome.
[Sources]
- https://doi.org/10.1093/humrep/deag096
- https://www.asrm.org/practice-guidance/practice-committee-documents/intracytoplasmic-sperm-injection-for-nonmale-factor-indications-a-committee-opinion-2026/
- https://doi.org/10.1093/hropen/hox011

## Slide 18
Beginner knowledge check
Answer aloud or use as a supervised discussion prompt
18
1
Which four interfaces must be controlled?
2
Why is an acellular droplet used before a live dish?
3
Name three findings that require a stop-use decision.
4
What is the difference between station readiness and an ICSI indication?
5
What must happen after a fault is corrected?
Continue learning
insideembryo.com
INSIDE EMBRYO BY AURION • EMBRYOLOGY & IVF LABORATORY SCIENCE
MANOJ KUMAR K • 18
Notes: Suggested answers: 1 vision, motion, pressure, environment. 2 to test function without risking reproductive material. 3 any three of drift, leak/bubble, damaged pipette, unstable temperature, vibration, identity gap. 4 technical release is not clinical indication. 5 recheck, document and release or escalate.
[Sources]
- https://doi.org/10.1093/humrep/deag096
- https://www.asrm.org/practice-guidance/practice-committee-documents/intracytoplasmic-sperm-injection-for-nonmale-factor-indications-a-committee-opinion-2026/

## Slide 19
Reference framework and final message
Use local regulation, manufacturer instructions and the laboratory SOP alongside these sources
19
Validated • aligned • responsive • environmentally stable • documented
That is a ready station—not a guaranteed outcome.
1. ESHRE. Recommendations on Good Practice in the IVF laboratory. Human Reproduction. 2026. doi:10.1093/humrep/deag096.
2. Kokkali G, et al. ESHRE good practice recommendations for polar body and embryo biopsy for PGT. Human Reproduction Open. 2020;2020(3):hoaa020.
3. Zegers-Hochschild F, et al. International Glossary on Infertility and Fertility Care, 2025. Human Reproduction. 2026. doi:10.1093/humrep/deag029.
4. ASRM Practice Committee. Intracytoplasmic sperm injection for nonmale factor indications: a committee opinion. 2026.
5. ESHRE Working Group on Embryologist Training Analysis. Training and competency assessment of Clinical Embryologists. Human Reproduction Open. 2023;2023(1):hoad001.
6. ESHRE SIG Embryology and Alpha Scientists. Vienna consensus on ART laboratory performance indicators. Human Reproduction Open. 2017;2017(2):hox011.
Educational use only. This deck does not replace supervised training, institutional protocols, professional guidance or patient-specific clinical advice.
INSIDE EMBRYO BY AURION • EMBRYOLOGY & IVF LABORATORY SCIENCE
MANOJ KUMAR K • 19
Notes: Close by repeating the release definition and the limits. Encourage learners to compare the sequence with their own laboratory's approved SOP under supervision.
[Sources]
- https://doi.org/10.1093/humrep/deag096
- https://doi.org/10.1093/hropen/hoaa020
- https://doi.org/10.1093/humrep/deag029
- https://doi.org/10.1093/hropen/hoad001
- https://doi.org/10.1093/hropen/hox011
- https://www.asrm.org/practice-guidance/practice-committee-documents/intracytoplasmic-sperm-injection-for-nonmale-factor-indications-a-committee-opinion-2026/